NephroTICK

Albustix Urine Dipstick Protein Classifier

# 1.  The Problem We Are Solving

Albustix dipstick tests are one of the most common ways to check for protein in urine — a key early warning sign for kidney disease. A nurse dips a strip into a urine sample, waits a few seconds, and compares the colour of the protein pad against a reference chart printed on the bottle. If the pad turns darker green, there is more protein.

The problem is that this visual comparison is highly subjective. Two clinicians looking at the same strip can disagree, and lighting conditions in a clinic can make the pad look different from the bottle chart. We wanted to replace that subjective step with a consistent, objective reading from a smartphone photograph.

Goal: Point your phone at the dipstick, press a button, and get a consistent protein reading — regardless of who is holding the phone or what the lighting looks like.

# 2.  Our Approach at a Glance

Rather than training a machine learning model from scratch (which would need thousands of labelled photos), we took a physics-based approach: measure the colour of the pad, compare it to a known colour scale, and report the closest match. This is exactly what a human expert does — we just do it with numbers instead of eyes.

The key insight is that if you photograph the reference colour chart alongside the strip in the same shot, both objects are lit by exactly the same light source. Any colour distortion caused by the phone camera or the room lighting affects both equally. We can therefore correct for it — and the comparison becomes much more reliable.

# 3.  What Happens When You Press Analyse

Here is the full pipeline, step by step, in plain English.

| Step | What happens and why |
| --- | --- |
| 1. Load the photo | The image is read from your upload or camera. Nothing is sent to a server — everything runs locally on the machine. |
| 2. White balance correction | Smartphone cameras add a warm (yellow) or cool (blue) tint depending on the room lighting. We correct for this by finding the neutral white and grey pixels in the background and scaling each colour channel so they look neutral. We deliberately ignore the green pad and the reference chart squares during this step — only the plain background is used — so coloured objects in frame (including latex gloves) do not skew the correction. |
| 3. Find the reference chart | We scan the image for all green-coloured blobs (using a colour range in HSV space that covers yellow-green through dark green). The reference chart on the bottle appears as six squares arranged in a tight column. We detect this column by looking for six blobs that all share nearly the same horizontal position — a much more reliable rule than just "find the biggest blobs". |
| 4. Find the protein pad | The protein pad sits on its own, away from the reference chart column. We identify it as the most isolated green blob — the one that is furthest from all the others. This works even when the chart and strip are both in frame. |
| 5. In-photo colour calibration | This is the most important step. We measure the colour of each of the six reference squares in the photo (under the current lighting). We then fit a mathematical correction that maps those measured colours to their known standard values. Applying this correction to the pad colour brings it into the same colour space as our reference database — regardless of whether the photo was taken under a hospital fluorescent light, a desk lamp, or in daylight. |
| 6. Extract pad colour | We average the corrected colours of the pixels inside the detected pad region, discarding the brightest and darkest 10% to reduce the effect of reflective highlights or shadows. |
| 7. Classify | We compare the corrected pad colour against our reference scale (NEG, Trace, 1+, 2+, 3+, 4+) using CIE Delta-E — the international standard for measuring how different two colours look to the human eye. The closest match within a confidence margin is reported. If the pad falls between two bands, the result is reported as "TRACE" and flagged for manual verification. |
| 8. Report | The result is displayed with the protein level, a clinical description, a confidence score (Delta-E), and a flag if clinician review is recommended (2+ and above, or ambiguous readings). |

# 4.  Why Calibration Matters

The single biggest source of error in colour-based dipstick reading is lighting variation. A pad that reads as Trace under a warm incandescent bulb can look like 1+ under a cool LED — even though it is the same strip.

Our solution is to always photograph the bottle reference chart alongside the strip in the same shot. The chart is printed with the standard colours for each protein level. Because the chart and the pad are illuminated by exactly the same light, any colour shift affects both equally. We measure the shift from the chart and reverse it before comparing the pad.

Technically, this is done with a least-squares affine transform in CIE L*a*b* colour space — a linear mapping fitted from six reference-to-standard colour pairs that corrects for differences in brightness, warmth, and hue simultaneously.

Analogy: Imagine you are trying to match paint colours but your screen has a yellow tint. If you put a known colour swatch on the screen next to the paint chip, you can see exactly how the screen is distorting things and correct for it. That is precisely what we do with the reference chart.

# 5.  How the Classification Decision Is Made

After calibration, every protein band on the reference scale has a known position in colour space. We measure the distance from the pad to each band using Delta-E (think of it as a ruler in colour space). The closest band wins — subject to a confidence check.

Each band is given a confidence margin equal to 75% of its distance to its nearest neighbour on the scale. If the pad's distance to the closest band is within that margin, we report that band. If it falls between bands — too far from any one of them — we report TRACE (ambiguous reading, verify manually).

| Result | Meaning |
| --- | --- |
| NEG | No protein detected. Normal. |
| SP/TR | Trace protein. Within normal limits. No action needed. |
| TRACE | Reading falls between two bands. Verify manually. Flagged for clinician. |
| 1+ | 30 mg/dL — mild proteinuria. Note and monitor. |
| 2+ | 100 mg/dL — moderate proteinuria. Flagged for clinician. |
| 3+ | 300 mg/dL — significant proteinuria. Flagged for clinician. |
| 4+ | 2000 mg/dL — severe proteinuria. Flagged for clinician. |

# 6.  Performance

We validated the model on 28 photographs of Trace (SP/TR) dipstick readings, taken following our standard protocol (strip and reference chart in the same shot). Each photo was taken under real-world conditions — different lighting, different camera angles, one photo even had a blue latex glove in the background.

| Stage | Accuracy on 28 Trace samples |
| --- | --- |
| Before in-photo calibration (IMG_8050 reference only) | 23 / 28 = 82% |
| After neutral-pixel white balance fix | 24 / 28 = 86% |
| With full in-photo calibration (affine warp) | 28 / 28 = 100% |

Important caveat: The 100% figure is on the dataset used during development. A held-out validation on new samples (different session, different concentrations) will be conducted when additional data becomes available.

The model also correctly classifies known 1+ samples in our original validation set, and correctly flags 2+ and above for clinician review.

# 7.  Known Limitations

- Single test kit only: calibrated for Albustix (Siemens). Other brands (e.g. Roche Combur) have different colour scales and are not supported.

- Reference chart must be in frame: if the chart is not photographed alongside the strip, the model falls back to a fixed reference and accuracy drops to ~86%.

- Single analyte: protein only. Other dipstick parameters (glucose, pH, blood, etc.) are not read.

- No training data generalisation test yet: we have not yet tested on photos from a different session, phone, or operator. This is planned for the next validation round.

- Not for clinical diagnosis: this is a decision-support tool. Any flagged result should be confirmed by a trained clinician.

# 8.  Technology Used

| Component | Technology |
| --- | --- |
| Image processing | OpenCV (Python) |
| Colour science | CIE L*a*b* colour space, CIE76 Delta-E |
| Colour calibration | Least-squares affine transform (NumPy) |
| Blob detection | HSV masking + contour analysis |
| Web interface | Streamlit (Python) |
| Platform target | iOS and Android (via browser or wrapped app) |
| Test strip | Siemens Albustix (protein-only) |